RANCANG BANGUN ALAT UKUR SUHU TINGGI DENGAN TERMOKOPEL Ni-Cu

Oleh:

Lino Hugun Saputra

11306141022

ABSTRAK
Penelitian ini bertujuan untuk (1) menyelidiki hubungan antara gradien suhu dangaya elektromotif (emf) pada sambungan Ni-Cu yang digunakan sebagai termokopel, (2) menyelidiki bentuk rangkaian listrik yang diperlukan untuk dapat mengkonversi dari input suhu menjadi tegangan, dan (3) menyelidiki karakteristik dari alat ukur yang terdiri dari sensitivitas, error, dan jangkauannya.

Penelitian ini dimulai dengan menyambung kombinasi plat nikel dan tembaga. Sambunganini dimasukkan kedalam cairan yang bersuhu tinggi dan suhu rendah. Sambungan tersebut dikarakteristik nilai keluaran tegangan termoelektrik yang diperoleh terhadap suhu. Keluaran tegangan termokopel yang diperoleh diperkuat menggunakan penguat diferensial dan mengkarakterisasi nilai perubahan tegangan terhadap suhunya. Hasil penguatan dihubungkan ke display berupa vu meter dan mengkalibrasinya menggunakan termometer. Data yang diperoleh digunakan untuk mencari nilai sensitivitas, error, jangkauan, dan keakuratan alat ukur.

Hasil penelitian menunjukkan (1) hubungan antara tegangan pada sensor termokopel dan suhu berbentuk polynomial, (2) penggunaan penguat differensial yang digunakaan sebagai penguat sinyal pada alat ukur suhu ini cukup efektif untuk sensor yang memiliki dua input tegangan, dan hubungan antara tegangan hasil penguatan dan suhu memenuhi persamaan grafik polynomial, serta (3) sensitivitas dari sensor termokopel yang dibuat adalah sebesar 0,05 mV/⁰C, danbesarerrorhisteresissetelahdilakukanperataanyakni 0,29 mV.
Kata-kata kunci:logamnikeldantembaga, tegangan termoelektrik, termokopel, penguat diferensial, vu meter.

CONSTRUCTION OF HIGH LEVEL TEMPERATURE MEASUREMENT WITH SENSOR TERMOCOUPLE Ni-Cu
By:

Lino Hugun Saputra

11306141022

ABSTRACT
The purpose of this research are (1) to investigate the relation between temperature and thermoelectric voltage from sensor termocouple Ni-Cu, (2) to investigate the electric combination in order to convert the voltage input from sensor termocouple become temperature,and (3) to investigate the characteristic of the device as sensitivity, error hysterisis, and range.
The first step in this research is tacked on the metal sheet which is nickel and copper. The junction was put into a liquid that has a difference temperature (hot themperature and cold themperature). Then from this junction we get the characteristic of thermoelectric voltage towards temperature. The output voltage of termocouple then reinforced with differential amplifier and find the characteristic building on alteration of voltage towards temperature. The voltage output of Op-Amp then connected with  vu meter as a display then calibrated with termometer. The data from this measurment then use to find sensitivity, error, range, and accuration of device.

The result of this researh are (1) the relation betwen temperature and thermoelectric voltage is polynomial, (2) the differential amplifier used in this device is effective enaugh for sensor which has two voltage input, and the relation between output voltage of Op-Amp and temperature is polynomial, and (3) the sensitivity of this device is 0,05 mV/⁰C, and the hysterisis error is 0,29 mV.

Keywords: nickel and copper metals, thermoelectric voltage, termocouple, differential amplifier, vu meter.
